EMERGENCY DEACTIVATION OF ICD/CRTD THERAPIES

Magnet application should be reserved for use in an EMERGENCY and only used until a Cardiac Device Physiologist can program the device off (if appropriate)

Device Physiologist ext 53851 or bleep 3031 ASAP (Mon – Fri 9am - 4pm)

For patients with ICD/CRTDs consideration may be given to positioning a clinical magnet over the implant site to inhibit inappropriate/unnecessary shock delivery (e.g therapy caused by noise detection or fast AF etc)

Please note the following:

- Inhibition of shock delivery will only be effective during magnet placement if   the magnet is properly secured to the patient! (use Micropore tape or equivalent)

- Any subsequent VT/VF will need to be treated using external defibrillation equipment, consider connecting the patient to an external defibrillator using remote pads in the AP position. In the event of a life-threatening arrhythmia, conventional advanced life support procedures should be followed

- Programmed bradycardia pacing will continue during magnet application of ICD/CRTD device